https://lucid.app/lucidchart/253e77ca-8b81-4c0a-8e63-2cc6b063a174/edit?docId=253e77ca-8b81-4c0a-8e63-2cc6b063a174&shared=true&page=HWEp-vi-RSFO#

Master Requirements Document 1. System Overview

The project is a portable waveform generator with FPGA-based control and app-driven interface.
 It must provide high-resolution, configurable signals across a range of frequencies, support standard lab equipment compatibility, and balance performance, portability, and usability.

System Objectives

- Generate high-quality arbitrary waveforms with 16-bit resolution.

- Provide portable and user-friendly operation with app-based control (ESP32).

- Ensure compatibility with standard lab measurement equipment (e.g., 50 Ω loads, BNC jacks).

- Demonstrate core capabilities in a 5-minute demo (signal generation, load driving, waveform control).

Key Design Constraints

- Performance Limits

  - Sampling rate: ~1 MHz max (1M samples/sec).

  - Nyquist compliance: ≥2× rate required for 500 kHz signals (avoid aliasing).

  - Output frequency: 100 kHz max usable waveform frequency.

- Resolution & Output Specs

  - 16-bit DACs (2 channels, parallel-based).

  - Output voltage: ±5 V.

  - Must drive 50 Ω load.

- Form Factor & Portability

  - Device must fit into a “nice enclosure.”

  - Table-top form factor.

  - Thermal management (heatsinks/grounding for DC/DC converters)

- Power

  - 12V DC Wall Plug

  - Must handle subsystem rails (FPGA, MCU, DAC, Op Amps, SRAM, etc.).

- Interface

  - No onboard interface required → App-based UI preferred.

  - Must allow waveform preview/control via phone app.

  - Optional: local display for debugging.

Required Features

- Signal Generation

  - Waveforms: sine, sinc, triangle, trapezoid, invert.

  - User-defined /  Optional: drawable waveforms.

  - Memory-backed waveform playback (FPGA + SRAM).

- FPGA Functionality

  - Index counter for waveform stepping.

  - RAM integration.

  - Load/playback control logic.

  - Configurable playback speed.

- Connectivity & App

  - App controls waveform generation, parameters, and playback.

  - Wireless communication via ESP32..

- Enclosure & I/O

  - BNC jacks for output.

  - Grounded, shielded enclosure.

  - Compact, user-friendly design.

Specifications Summary

| Parameter | Requirement |
| --- | --- |
| DAC Resolution | 16-bit, 2 channels |
| Max Sampling Rate | Minimum ~1 MHz (1M samples/sec) |
| Max Output Frequency | Minimum: ~100 kHz |
| Output Voltage Range | ±5 V |
| Output Load | 50 Ω |
| Coupling | Selectable AC/DC |
| Waveforms Supported | sine, sinc, triangle, trapezoid, invert, custom |
| Memory | SRAM for waveform storage |
| Controller | FPGA (index counter, RAM, playback) |
| User Interface | IOS App (ESP32-based) |
| Ports | 2 BNC connectors (to connect to oscilloscope) |
| Power Options | DC 12V Wall Power |

7. Demo Deliverables (Dec. 2)

- Demonstrate 16-bit waveform generation at ±5 V.

- Show multiple waveforms (sine, sinc, triangle, trapezoid, invert, etc).

- Control waveform parameters via app (ESP32).

- Prove ability to drive a 50 Ω load.

- Demonstrate clock speed control and waveform playback from memory.